## Comments
Draft Identifier: | AS 23907 | Sharps injury protection — Requirements and test methods — Sharps containers (ISO 23907-2:2019, MOD)
Section | Type of Comment | Comment | Proposed Change
sec_preface | Technical | (a) (v) Requiring all sharps containers to have a temporary and permanent closure: it is WMRR’s understanding that based on the current market and types of containers currently available, this requirement could have a detrimental effect on market dynamics, with a significant cost impost borne on users of non-reusable containers. | It is imperative that Standards Australia commissions an independent party to undertake a full analysis of the current range of containers used in Australia to determine the implications of this new requirement on users and suppliers. Following this analysis, the Waste Management and Resource Recovery Association of Australia's (WMRR) Biohazardous Waste Industry (BWI) division proposes that the requirement is reviewed and a further comment process undertaken before the it is finalised.
sec_preface | Technical | (a) (vi) Requiring all sharps containers to have penetration requirement of 20N: while WMRR agrees with the nominal 20N requirement for reusable contrainers as it appears to work and no change would impact on current manufacturers or market dynamics, WMRR queries if a change is necessary or appropriate for non-reusable, single-use containers, which would impact current suppliers and market dynamics. | WMRR recommends that the requirement for non-reusable containers remains at 12.5N and that any proposal to increase this requirement should be reviewed, qualified, and determined by an independent technical advisor.
sec_preface | Technical | N.A. | (a) (vii) Redefining fill line capacity: WMRR recommends that the fill line capacity be set at 85% or less.
sec_preface | Technical | (b) (iii) Decreasing the fluid and mass in the drop test: the ratios used are incorrect. Where it was previously less than 60ml (for single-use containers), it is now proposed that 1% of 86% volume is only less for sharps containers under seven (7) litres but could be as much as 510ml for a 60-litre container, effectively increasing the fluid volume by 10 times. | N.A.
sec_preface | Technical | (b) (i) Enable single-use and reusable sharps containers to be included in one document: while WMRR appreciates the intent behind publishing a single document for all containers, the difference in usage of single-use vs. reusable containers must both be clearly understood and articulated in the standards, including greater clarity in the differences in physical parameters. | WMRR proposes that at a minimum, a table must be included to compare the similarities and differences between containers. WMRR also believes that an independent review must be undertaken to determine the relevance and suitability of any proposed change that seeks to align the two (2) categories, including but not limited to: 1. The current safety concerns surrounding the status quo requirements. 2. The selection of proposed requirements and the supporting evidence. 3. An independent market impact analysis that outlines the cost benefit of proposed changes and the likely lessening of competition or increase in costs to consumers.
sec_intro | Editorial | N.A. | 2. Delete second paragraph: Sharps containers can be either single-use or reusable. This document covers reusable sharps containers: WMRR recommends the deletion of only “this document covers reusable sharps containers”, as the first sentence of this paragraph is accurate and is important for abundance of clarity.
sec_3 | Editorial | 3.2 Clean - visually free of soil, debris and organic matter: this definition needs to be improved. For one, being visually free of soil, debris, and organic matter does not necessarily constitute cleanliness. | The definition also needs to extend the scope such that the container can be considered “free from pathogens and bacteria” and the container can be verified that it has undergone a qualified decontamination process; furthermore, this process must have its thresholds defined in the standards and WMRR recommends seeking independent validation from infection control stakeholders. Specifically, WMRR recommends that Appendix A - MIcrobiological validation, is included as part of the standards (not as an appendix) as it is vital for the determination of "cleanliness", which itself is a critical feature of reusable containers. It must be noted that 3.2 relates only to reusable containers as single-use containers would be clean prior to use and would only be used once.
sec_3 | Technical | 3.17 Representative sharps - Representative quantities of unused 1ml to 60ml syringes-needles in the ratio of 70% safety devices with safety mechanism activated and 30% syringe-needles with no safety device: this needs to be reviewed as verification for Australia is not fully represented. Needles with safety devices are not commonplace in Australia and the references used relate to the US. | Seek feedback from Australian clinical references for what a suitable "representative" batch would include
sec_3 | General | 3.20 Secondary stabilizer - attachment or design feature intended to provide extra stability and prevent the device from toppling over: a container must be compliant with all safety requirements on its own or any test will be redundant; WMRR does not agree with creating safety features that provides opportunities for users/suppliers to evade the requirements. Additionally, any container must be safe on its own merit as there may be situations where these containers cannot be placed on a stabiliser. | WMRR proposes that if this requirement is kept, there must be a guarantee that the container will attached to the secondary design (the stabiliser) and that a fail safe device is required to ensure that the container cannot be used unless it is attached to the secondary stabiliser.
sec_4 | Editorial | 4.1 General – national variations: the proposed paragraph refers to a range of different subtypes of sharps with no clear guidance on what constitutes each, and therefore how and why each require a different colour container. Pharmaceutical sharps is one (1) example where a very broad range of opinions exist as to what level of contaimination is required to justify incineration. Further, WMRR is seeking clarification on the colour codes currently listed in AS4031, specifically that the container intended for the collection of radioactive waste should be predominantly red. WMRR notes that red has already been removed from the colour chart. | There is an opportunity to improve labelling and definitions in the standards. As such, WMRR suggests that a clear definition of each waste type and the situation where a sharp warrants a separate waste stream apart from the standard biohazardous sharps. Clarity is required to ensure that only objects defined as sharps are placed in sharps containers There is also no need to reference incineration in (c).
sec_4 | Editorial | 4.2.1 (2): this inclusion is not required as it is irrelevant to both the scope and to safe handling; this inclusion relates to legislation where it should remain. | N.A.
sec_4.2.1 | Technical | Point 3: both (a) and (b) will impact the competitive advantage of suppliers. WMRR is seeking a rationale for why these requirements are exclusive to only 30L containers | If containers are designed to be stacked, they must be designed to pass the topple test and have a label indicating that it is safe to do so. Similarly, a label for containers not meant to be stacked must be adhered to these containers. Broadly, WMRR proposes removing both (a) and (b) as these requirements are propriety features and out of scope of these standards.
sec_4.2.2 | Technical | 4.2.2 Container stability and 6.1.2: Clarity is sought on what “most adverse position” entails, including a definition and/or parameters. The exclusion under “NOTE” is too broad and could allow a number of variants to be exempt from a legitimate test as users may not be aware that the exemption only applies to wall-mounted and pocket collectors. Further, section seven (7) is out of scope for this document and would be adequately covered under WH&S policies. | If a container is designed and approved for wall mounts only, then labelling requirements are necessary, in line with 8.1 – Labelling and marking, and 8.2 – Instructions for use.
sec_4.2.4.1 | Editorial | N.A. | The term “reusable” should not be deleted. Further indicators need to also be developed and included for “reusable”.
sec_4.2.4.3 | Editorial | Typically, if a container has a permanent closure, it is not identified or deemed as reusable. The issue at hand is the lack of definition around closures. | A review is recommended that must conclude with clear definitions of each type of closure and how that they would apply to either single use, reusable or both. WMRR would suggest that there are at least three (3) types of closures - permanent (in the case of a single use container - NEVER to be opened again), temporary (both), and temporarily permanent (in the case of a reusable container). Clarity is also required around whether these requirements solely relate to reusable containers. Following this review, WMRR recommends that this section is re-written, specically that it is broken up into two (2) sections - one for reusable containers and the other for single-use containers - with each section detailing clearly the requirements related to the specified container. Additionally, noting that this document relates to all sharps containers, WMRR queries how one would temporarily close and re-open a screw-top container with only one (1) hand.
sec_4.2.5 | Technical | As noted above, WMRR queries why the penetration test for single-use containers has increased from 12.5N to 20N and requests the release of metrics and analyses behind this proposed change, as well as all evidence to show that (a) the current standard is unsafe, including data, if any, to show that injuries have occured as a direct result of the current standard, and (b) evidence that there are suppliers who can adhere to this as this new requirement may negatively impact market competitiveness. | WMRR suggests that this is replaced with “for reusable containers and 12.5N for single-use containers”. If the data and evidence requested are unavailable, independent validation of this requirement is necessary.
sec_4.2.7 | General | Clarity is sought around the temporary closure mechanics for screw-top sharps containers; an outline of what a temporary closure is or isn't - as a screw top closure may not be temporary - is required. | N.A.
sec_4.5.2 | General | WMRR queries the relevance of this statement, noting that being considered “desirable” does not make it a requirement. | N.A.
sec_6.4.2.2 | Technical | N.A. | A review is required, and it is suggested that this be changed to 0.85% of the maximum volume. To set the same area for a 510ml container as a 60L container is an unrealistic requirement for sharps containers.
sec_A | Editorial | N.A. | As per comments in section 3.2, it is proposed that Appendix A is moved into the main body of the standards.
sec_bibl | General | WMRR queries why the draft standards have referenced US legislation and transport regulations that do not apply to Australia, which has its own federal and jurisdictional regulations. | N.A.

## Instructions
Instructions in Using This Template Correctly
1. Enter your comment details starting from row 3 downwards and in columns A, B, C, and D only.
2. Do not change anything in rows 1 and 2. This will cause your bulk upload to have errors and your comments will not be saved into the draft.
3. Any details that you will enter outside of columns A to D will be disregarded.
4. For your entries in columns A and B, you can only select values from the dropdown lists. Any entry not found in the lists will cause an error.
5. You are allowed to change the filename of this template.
6. For each row where you have an entry, you must fill out columns A to D as they are mandatory.
7. Make sure that you click on F7 to perform a Spell check and correct any errors before saving the file.
8. For further assistance, click the 'Contact Standards Australia' link in the footer section.

## Section (hidden)
sec_titlepage
sec_preface
sec_foreword
sec_intro
sec_1
sec_2
sec_3
sec_4
sec_4.1
sec_4.2
sec_4.2.1
sec_4.2.2
sec_4.2.3
sec_4.2.4
sec_4.2.4.1
sec_4.2.4.2
sec_4.2.4.3
sec_4.2.5
sec_4.2.6
sec_4.2.7
sec_4.2.8
sec_4.3
sec_4.4
sec_4.5
sec_4.5.1
sec_4.5.2
sec_5
sec_5.1
sec_5.2
sec_5.3
sec_5.4
sec_5.5
sec_6
sec_6.1
sec_6.1.1
sec_6.1.2
sec_6.2
sec_6.2.1
sec_6.2.2
sec_6.2.3
sec_6.2.4
sec_6.3
sec_6.3.1
sec_6.3.1.1
sec_6.3.1.2
sec_6.3.1.3
sec_6.3.1.4
sec_6.3.2
sec_6.3.2.1
sec_6.3.2.2
sec_6.3.2.3
sec_6.3.2.4
sec_6.3.2.5
sec_6.4
sec_6.4.1
sec_6.4.1.1
sec_6.4.1.2
sec_6.4.1.3
sec_6.4.2
sec_6.4.2.1
sec_6.4.2.2
sec_6.4.2.3
sec_6.4.2.4
sec_6.5
sec_6.5.1
sec_6.5.2
sec_6.5.2.1
sec_6.5.2.2
sec_6.5.2.3
sec_6.5.2.4
sec_7
sec_8
sec_8.1
sec_8.2
sec_A
sec_A.1
sec_A.2
sec_A.2.1
sec_A.2.2
sec_A.2.3
sec_A.3
sec_A.3.1
sec_A.3.2
sec_A.3.3
sec_A.3.4
sec_A.3.5
sec_A.3.6
sec_bibl

## TypeofComment (hidden)
General
Editorial
Technical